8/10/2023

- BradyParameter: program takes too long, needs to assign more response time before clicking retrieve button

- BatteryLog:

- TachyLog:

- ResetLog:

- ChargeLog:

- Manufacturing:

Total Running time: 25mins 40s

10 failed, 32 passed

8/11/2023

- Brady has the issue of overlap with tachy detection

- Lower Rate limit with max Sensor Rate

Third Large Scale Unit Test

Probably the cost of transmission layer failure. All other test cases pass.